## X tube QA prot
Radiological Protection and Quality Assurance Report
Hospital/Clinic: |  |  |  |  | Unit type :
Department : |  |  |  |  | Generator/Manuf.:
Type:
Room/Unit: |  |  |  |  | X-Tube / Type:
Ser.No.
Date of QA survey: |  |  |  |  | Reference no.:
SUMMARY AND RECOMMENDATIONS
Report compiled by:
Dr. S.D.Tabakov
RESULTS
Parameters |  |  |  | Expected |  | Measured |  | Assessment
Radiography Mode
Beam/LBD Alignement |  | displacement (cm) |  | <1 |  | <1 |  | Accept
Bucky centering |  | displacement (cm) |  | <1 |  | <1
Focus size (mm) |  | Broad Focus-(BF) |  | 1 |  | 1X1 |  | Accept
|  | Fine Focus-(FF) |  | 0.6 |  | 0.6x0.6 |  | Accept
O'put consistency(BF) |  | (%) inconsist. |  | <5 |  | 3.2416 |  | Normal
O'put var'n with mA(BF) |  | (%) variation |  | <10 |  | 9.4884 |  | Normal
O'put var'n with kV (BF) |  | Linearity (%) |  |  |  | 4.1512 |  | Normal | Linearity (stdev O128;O130) |  | 9.4592e-07
Specific O'put-80kV(BF) |  | (uGy/mAs) @ 1m |  |  |  | 42.7395 |  | Normal
Lin.Grad.with kV^2(BF) |  | (mGy/mAs/kV^2) |  |  |  | 6.2665e-06 | 15.095 |  | Grad' (mGy/mAs)/kV^2 (LINEST) |  |  | 6.3232e-06
O'put var'n with kV (FF) |  | Linearity (%) |  |  |  | 8.4402 |  | Normal | Linearity (stdev O133:o135) |  | 3.3743e-06
Specific O'put-80kV(FF) |  | (uGy/mAs) @ 1m |  |  |  | 57.4721 |  | Normal
Lin.Grad.with kV^2(FF) |  | (mGy/mAs/kV^2) |  |  |  | 8.4827e-06 | 39.7781 |  | Grad' (mGy/mAs)/kV^2 (LINEST) |  |  | 7.8569e-06
kVp consistency (BF) |  | (%) inconsist. |  | <5 |  | 0.60423 |  | Good
kVp accuracy (BF) |  | (%) inaccuracy |  | -10<A<10 |  | 3.3958 |  | Good
kVp accuracy (FF) |  | (%) inaccuracy |  | -10<A<10 |  | 0.74315 |  | Good
|  | St.dev. |  |  |  | 2.3082
kVp var'n with mA (BF) |  | (%) variation |  | <10 |  | 4.2503 |  | Good
Ripple at ~100 kV(BF) |  | +/- kVp ampl. |  |  |  | 5 |  | Accept
Timer consistency (BF) |  | (%) inconsist. |  | <5 |  | 0 |  | Good
Timer accuracy (BF) |  | (%) inaccuracy |  | -10<A<10 |  | -1.5833 |  | Good
|  | St.dev |  |  |  | 3.9555
First Half Value Layer |  | (mm of Al eq') |  |  |  | 2.7041 |  | Accept
Inferred Total Filtration |  | (mm of Al eq') |  | >2.5 |  | 2.6 |  | Accept
Labelling | [7 empty] | Accept
Stated filtration |  | (mm of Al eq') |  | 2.5
Stated Total Filtration |  | (mm of Al eq')
Calcul. max. leakage |  | (mGy/h)@1m |  | <1 |  |  |  | Accept
X-RAY GENERATOR AND TUBE MEASUREMENTS
These measurements were made with a keithley kVp divider s/n xxxx, digital storage oscilloscope
Gould 450 s/n zzzz, and an MDH electrometer s/n yyyy with 6 cc chamber .
[7 empty] | FDD(cm)= | 100
Focus | Set kV | Set mA | Set Time | Set mAs | Meas kV | Meas T | Meas exp | Air kerma |  | Meas' | mGy/mAs | mGy/mAs | For lin' | %age kV | %age T | Meas' | mGy/mAs | mGy/mAs
| (kV) | (mA) | (ms) | (mAs) | (kV) | (ms) | (mR) | (mGy) | Meas kV | kV^2 | mGy/mAs | per kV^2 | grad' | Accuracy | Accuracy | kV^2 | Broad focus | Fine focus
B | 60 | 200 | 100 | 20 | 62 | 101 | 59 | 0.51271 | 62 | 3844 | 0.025635 | 6.669e-06 | 5.6364e-06 | 3.3333 |  | 3844 | 0.025635
B | 80 | 200 | 100 | 20 | 83 | 105 | 98.5 | 0.85596 | 83 | 6889 | 0.042798 | 6.2125e-06 | 5.8088e-06 | 3.75 |  | 6889 | 0.042798
B | 100 | 200 | 100 | 20 | 104 | 103 | 151 | 1.3122 | 104 | 10816 | 0.06561 | 6.066e-06 | 7.3542e-06 | 4 |  | 10816 | 0.06561
B | 120 | 200 | 100 | 20 | 123 | 103 | 224 | 1.9466 | 123 | 15129 | 0.097328 | 6.4332e-06 |  | 2.5 |  | 15129 | 0.097328
F | 50 | 100 | 200 | 20 | 50 | 202 | 46 | 0.39974 | 50 | 2500 | 0.019987 | 7.9948e-06 |  | 0 |  | 2500 |  | 0.019987
F | 70 | 100 | 200 | 20 | 73 | 205 | 101 | 0.87769 | 73 | 5329 | 0.043884 | 8.235e-06 | 8.4473e-06 | 4.2857 |  | 5329 |  | 0.043884
F | 90 | 100 | 200 | 20 | 88 | 202 | 167 | 1.4512 | 88 | 7744 | 0.072562 | 9.37e-06 | 1.1875e-05 | -2.2222 |  | 7744 |  | 0.072562
F | 110 | 100 | 200 | 20 | 111 | 204 | 221 | 1.9205 | 111 | 12321 | 0.096024 | 7.7936e-06 | 5.1263e-06 | 0.90909 |  | 12321 |  | 0.096024
B | 80 | 200 | 100 | 20 | 82 | 105 | 104 | 0.90376 | 82 | 6724 | 0.045188 | 6.7204e-06
B | 80 | 200 | 100 | 20 | 83 | 105 | 106 | 0.92114 | 83 | 6889 | 0.046057 | 6.6856e-06
B | 80 | 200 | 100 | 20 | 83 | 105 | 105 | 0.91245 | 83 | 6889 | 0.045622 | 6.6225e-06
B | 80 | 25 | 100 | 2.5 | 83 | 107 | 18.4 | 0.1599 | 83 | 6889 | 0.063958 | 9.2841e-06 |  |  | 7
B | 80 | 300 | 100 | 30 | 84.7 | 105 | 198 | 1.7206 | 84.7 | 7174.09 | 0.057354 | 7.9946e-06 |  |  | 5
B | 80 | 500 | 100 | 50 | 90 | 103 | 305 | 2.6504 | 90 | 8100 | 0.053009 | 6.5443e-06 |  |  | 3
B | 80 | 200 | 20 | 4 | 83 | 19 | 20.5 | 0.17814 | 83 | 6889 | 0.044536 | 6.4648e-06 |  |  | -5
B | 80 | 200 | 400 | 80 | 83 | 411 | 440 | 3.8236 | 83 | 6889 | 0.047795 | 6.9379e-06 |  |  | 2.75
B | 80 | 200 | 800 | 160 | 83 | 780 | 814 | 7.0737 | 83 | 6889 | 0.04421 | 6.4175e-06 |  |  | -2.5
Half Value Layer Measurements are shown below
B+0mm Al | 80 | 200 | 100 | 20 |  |  | 114 | 0.99066 | HVL mm of Al | % trans |  |  |  |  | #VALUE! | HVL | % trans
B+0mm Al | 80 | 200 | 100 | 20 |  |  | 110 | 0.9559 | 1 | 100 | [5 empty] | 0 | 100
B+1mm Al | 80 | 200 | 100 | 20 |  |  | 83.5 | 0.72562 | 2 | 74.5536 | [5 empty] | 1 | 74.5536
B+2mm Al | 80 | 200 | 100 | 20 |  |  | 67.1 | 0.5831 | 3 | 59.9107 | [5 empty] | 2 | 59.9107
B+3mm Al | 80 | 200 | 100 | 20 |  |  | 50.9 | 0.44232 | 4 | 45.4464 | [5 empty] | 3 | 45.4464
B+4mm Al | 80 | 200 | 100 | 20 |  |  |  | #VALUE! | 5 | #VALUE! | [5 empty] | 4 | #VALUE!
|  |  |  | #VALUE! |  |  |  | #VALUE! | [6 empty] | #VALUE!
|  |  |  | #VALUE! |  |  |  | #VALUE! | [6 empty] | #VALUE!
The X-ray generator and tube perform well . However, the X-ray beam is about 1cm displaced from the light beam of LBD in longitudinal direction. We would advise you the bring this to the attention of the service engineer on his next routine visit.