## Slide 1: What do we mean by quality and what is the Role of the Medical Physicist ?
Assoc Prof Paddy Gilligan, Dublin Ireland ( with assistance from Harry Delis and Annalisa Trianni )

## Slide 2: Quality
Conformance to requirements

## Slide 3: What is Quality ?
A degree of excellence
Conformance to requirements
Totality of characteristics which act to satisfy a need
Fitness for use
Fitness for purpose
Freedom from defects
Delighting customers
“… the degree to which a set of inherent characteristics
fulfils a need or expectation that is stated, generally
implied or obligatory” [ISO 9000]

## Slide 4: “Primum non nocere”
First, do no harm - The most important rule of healthcare !!
Especially in radiation medicine, either for diagnosis or for treatment, this is even more evident as the “harm” associated with radiation is well-known.
Some tools that modern medicine has in order to monitor that radiation medicine is performed safely and effectively are:
  Quality Processes and Dosimetry

## Slide 5: Quality health care can be defined in many ways but there is growing acknowledgement that quality health services should be:
Effective – providing evidence-based healthcare services to those who need them;
Safe – avoiding harm to people for whom the care is intended; and
People-centred – providing care that responds to individual preferences, needs and values.
To realize the benefits of quality health care, health services must be:
Timely – reducing waiting times and sometimes harmful delays;
Equitable – providing care that does not vary in quality on account of gender, ethnicity, geographic location, and socio-economic status;
Integrated – providing care that makes available the full range of health services throughout the life course;
Efficient – maximizing the benefit of available resources and avoiding waste.

## Slide 6: Sustainable health care
Sustainability is defined by the World Wildlife Fund for Nature (WWF) as “an economic activity that meets the needs of the present generation without compromising the ability of future generations to meet their needs”.
A sustainable health care system is based on delivering high quality care and improved public health and wellbeing without exhausting natural resources or causing adverse ecological damage.

## Slide 7: Medical physicists apply physics to benefit patients and staff :
Relevant to the patients, demographics we serve
Accessible and affordable treatment and diagnostics which minimise
high standard of training availability
Able to sustain shocks in the modern world
  climate,
  war , cyber and physical
  unsteady electrical supply
  Financial shocks
Minimum harm to the environment and public health e.g radioactive waste , use of highly toxic agents in production,
Optimisation and audit

## Slide 8: We are all part of eco system
Society
Healthcare system
Hospital
Medical Physicist
World

## Slide 9: What is the mission statement of your hospital ?

## Slide 10: Types of fluoroscopic equipment:
Mini c-arms
C-arms with image intensifiers
3d flat panel detector
R/F fluoro rooms
C-arm angio, cardiovascular, radiology
CT fluoroscopy
O-arm fluoroscopy
Hybrid systems e.g. CT and C-arm
Imaging systems on linear accelerator

## Slide 11: why these devices are important
critical life saving equipment involved in complex procedures often outside radiology
Potential for highest patient and staff doses with tissue effects
  Keep your clinicians working and your patients from erythema
Expensive piece of equipment in constant state of evolution
Complex human machine interaction

## Slide 12: Exams using ionising radiation, Ireland 2009-2012

## Slide 13
Exams using ionising radiation, Ireland 2009-2012

## Slide 14: Occupational doses from interventional systems can be the highest Occupational doses in any hospital
Chart: Cardiologist Outside Apron Doses 2011, eye dose surrogate (mSv)
MPH | 6.5 | 3.6 | 7.1 | 22.8 | 24.4 | 2.5 | 4.5 | 11.7 | 38.7 | 44.4 | 36.2 | 26.3 | 37.6 | 0.1 | 1.2 | 5.9
MMUH | 0 | 9.28 | 0.25 | 0 | 0 | 1.88 | 4.08 | 22.8 | 11 | 0 | 3.9 | 0 | 15.12 | 0.65 | 0 | 0
Classification
Limit

## Slide 16: Associated radiation effects :absence of evidence versus evidence of absence

## Slide 17: 18% growth in fluoro systems
https://www.itnonline.com/article/fixed-digital-radiography-and-fluoroscopy-see-market-growth

## Slide 18
Challenges in the near future – XA ( courtesy Johan Sjoberg)
KS+KH pre 2016
Today: 17 labs
NKS + CHOPIN (KH)
>2016: ~ 35 labs
?
?
?
?
?
?

## Slide 19: How will we resource the medical physics needs ?
Workforce planning
Surveys EU rest
Training syllabus
Regulatory compliance
Provide a meaningful and impactive physics service aligned with hospital care system goals
Review this service
Value

## Slide 20: How to achieve Quality
“Quality does not happen by accident, it must be planned”
“Quality is everyone's responsibility”
Joseph Moses Juran
(1904 – 2008)
W. Edwards Deming
(1900 - 1993)

## Slide 22: Management System
Quality has to be achieved and maintained by means of an effective management system
“The term management system reflects and includes the concept of ‘quality control’ (controlling the quality of products) and its evolution through ‘quality assurance’ (the system for ensuring the quality of products) and ‘quality management system’ (the system for managing quality)”
2014

## Slide 23: Common terms
Summary of terms and definitions:
QM: An overall management system that includes establishing quality policies and quality objectives, and processes to achieve quality objectives through quality planning, QA, QC, and quality improvement.
QA: A component of QM focused on providing confidence that quality requirements will be fulfilled; it includes all activities (planned, systematic, and practice‐based activities) that demonstrate the level of quality achieved by the output of a process.
QC: A component of QM focused on the fulfilment of quality requirements; it includes activities that impose specific quality on a process; and entails the evaluation of actual operating performance characteristics of a device or system, comparing it to desired goals, and acting on the difference; QC works on the input to a process to ensure that important elements or parameters specific to the process are correct.

## Slide 24: EU BSS 59/2013
"quality assurance" means all those planned and systematic actions necessary to provide adequate assurance that a structure, system, component or procedure will perform satisfactorily in compliance with agreed standards. Quality control is a part of quality assurance;
"quality control" means the set of operations (programming, coordinating, implementing) intended to maintain or to improve quality. It includes monitoring, evaluation and maintenance at required levels of all characteristics of performance of equipment that can be defined, measured, and controlled;
Proactive process to prevent defects
Reactive process to detect defects

## Slide 25: Optimisation & clinical audit
Member States shall ensure that the optimisation includes the selection of equipment, the consistent production of adequate diagnostic information or therapeutic outcomes, the practical aspects of medical radiological procedures, quality assurance, and the assessment and evaluation of patient doses or the verification of administered

## Slide 26: medical physics expert /BSS EU 13/59
(a) optimisation of the radiation protection of patients and other individuals subject to medical exposure, including the application and use of diagnostic reference levels;
(b) the definition and performance of quality assurance of the medical radiological equipment;
(c) acceptance testing of medical radiological equipment;
(d) the preparation of technical specifications for medical radiological equipment and installation design;
(e) the surveillance of the medical radiological installations;
(f) the analysis of events involving, or potentially involving, accidental or unintended medical exposures;
(g) the selection of equipment required to perform radiation protection measurements;
(h) the training of practitioners and other staff in relevant aspects of radiation protection;
3. The medical physics expert shall, where appropriate, liaise with the radiation protection expert.

## Slide 27: Quality management system
Procedures
Documentation and document control
Equipment performance
Procurement
Environment
Responsibilities
Training & Personnel
Dose reference levels
Fault logs
Procedures
clinical audits
Incidents
Risk register
Knowledge sharing and training

## Slide 29: Standards
Clinical audit
Accreditation
Regulatory compliance and inspection

## Slide 30: What is a clinical audit ?
One of the key elements within the framework of quality is the comprehensive and critical review of the practice in a medical radiation facility

## Slide 31: What is a clinical audit ?
One of the key elements within the framework of quality is the comprehensive and critical review of the practice in a medical radiation facility
  Independent external evaluation  independent assessment of how actual clinical practice compares with good practice

## Slide 32: What is a clinical audit ?
One of the key elements within the framework of quality is the comprehensive and critical review of the practice in a medical radiation facility
  Independent external evaluation  independent assessment of how actual clinical practice compares with good practice
  Peer review evaluation of the quality of practice  any areas for improvement will be identified with a view to bring the services to the internationally accepted level

## Slide 33: What is a clinical audit ?
One of the key elements within the framework of quality is the comprehensive and critical review of the practice in a medical radiation facility
  Independent external evaluation  independent assessment of how actual clinical practice compares with good practice
  Peer review evaluation of the quality of practice  any areas for improvement will be identified with a view to bring the services to the internationally accepted level
Quality Improvement

## Slide 34: What is a clinical audit ?
Clinical Audit is “a systematic examination or review of medical radiological procedures which seeks to improve the quality and outcome of patient care through structured review, whereby medical radiological practices, procedures and results are examined against agreed standards for good medical radiological procedures, with modification of practices, where appropriate, and the application of new standards, if necessary”
Council Directive 2013/59/EURATOM

## Slide 35: What a clinical audit IS NOT
Clinical audits are NOT designed for:
  Regulatory purposes
    Not enforcing tool but of advice on quality improvement
  Investigation of accidents or reportable medical events
    A more focused audit is required
  Assessment for entry into cooperative clinical research studies

## Slide 36: What does a clinical audit cover?
A comprehensive clinical audit of diagnostic radiology practices consists of a review and evaluation of the quality of all elements involved in the practices:
  Staff
  Equipment
  Procedures, patient protection and safety
  Overall performance of the facility
  Interaction with external service providers

## Slide 37: Time staff dose and motion study Cardiology

## Slide 38: Accreditation
Accreditation is a formal, independent verification that a program or institution meets established quality standards and is competent to carry out specific conformity assessment tasks. Conformity assessment tasks may include, but are not limited to, testing, inspection, or certification

## Slide 39: Role of the physicist in life cycle of equipment

## Slide 40: Role of the medical physicist

## Slide 41
Think sustainable at all stages of the life cycle , from design , use to disposal
MPP’s potential roles and degrees of involvements in the stages / procedures within a medical device life cycle Responsible (r), accountable (a), consulted (c), or informed (i) Involvement: High (H), Medium (M) or Low (L) |  |  | 
Procurement: | Definition of requirements | r/a | M
 | b. Market analysis of potential technical solutions / vendors | r/a | H
 | c. Requirements / specifications, standardization | r/a | H
 | d. Decision criteria / quote evaluation | r/a | M
 | e. Business case / Budget | c/i | M
 | f. Price; TCO; cost effectiveness | c/i | M
 | g. Availability of upgrades / improvements | r/a | H
 | h. Maintenance options / contracts | a/c | M
 | i. Risk management | r | H
 | j. Training possibilities | a/c | H
 | k. Investment decision / placing order | c/i | L
Quality control / assurance | a. Working procedures | r | H
 | b. Assessments | r | H
 | c. Audits | r/a | H
Implementation / Commissioning | a. Acceptance | r/a | H
 | b. Application | a | M
Training | a. (Technical) application training | a | H
 | b. Maintenance training | a/c/i | M
Operation | a. Maintenance | c/i | L
 | b. Administration / management | a | L
 | c. Optimization and training | r | H
 | d. Updates and upgrades | r/a | H
Decommissioning | a. Discharge / disassemblance | a | M

## Slide 42: Types of testing

## Slide 43: Physics Role in quality assurance
On site as part of team
Visit out of hours
Remote monitoring

## Slide 44: Testing equipment does not become quality control unless results are acted upon ;

## Slide 45: Standards we use for equipment monitoring
RP 162,
IEC
NEMA
RP 91
XR 27
Published DRLS
EFOMP angio QC protocol ( in press)
EFOMP CBCT protocol

## Slide 46: Look around
...and update

## Slide 47: Look around

## Slide 48: Look around
Need for harmonization

## Slide 50: I do a lot of tests but they always pass, why should I bother ?
1-2 %

## Slide 51: EU 59/13
Accidental and unintended medical exposures are a source of continuing concern …
It is the role of the competent authority in radiation protection to address the prevention of accidental and unintended medical exposure and the follow-up in case of their occurrence…
In this respect, the role of quality assurance programmes, including a study of risks in radiotherapy, to avoid such incidents should be emphasised, and recording, reporting, analysis and corrective action should be required in such cases.

## Slide 52: Incident analysis teams :
The person in charge of the facility/department/staff where the incident took place;
The person who was exposed;
The person(s) acting as operator during the exposure;
The Radiation Protection Officer; and
The Radiation Protection Expert / Medical Physics expert.
In the case of an incident which occurred as a result of defective equipment the following people should also be involved in the investigation:
  The service engineer who examined the equipment following the incident; and
  The person who was responsible for quality assurance on the equipment.

## Slide 53: The incident reporting toolkit
What has happened the patient and their care
The guidelines ( read them)
Good safety culture, Active RSO & trained staff
Incident recording software
Dose tracking software
Quality structures for investigation/reporting
Medical physics expert with appropriate dose software
Resource to report and follow up on findings
Patient disclosure training
Records

## Slide 54: Every system is not always perfect but we can use a priori knowledge to mitigate future risks
Incidents, Near Misses, root cause analysis
Fault logs , heat maps, uptime
Service report
Service portals
Proactive monitoring alerts
Hazard alerts field safety notices
Recalls
FDA Maude Data base
Root cause analysis
Trends Parieto analysis
Newsgroups
Publications

## Slide 55: Estimation of effective and organ doses in patient undergoing interventional neuroradiology procedures: Preliminary results of a multicenter study
Andrea D’Alessio1, Barbara Cannillo1, Giulia Azzalin2, Roberta Matheoud1, Giuseppe Guzzardi3, Andrea Giorgianni4, Marco Brambilla1, Sabina Strocchi5
(1) Medical Physics Unit - AOU Maggiore della Carità - Novara - Italy
(2) School of Medicine - Università del Piemonte Orientale A. Avogadro - Novara - Italy
(3) Radiology Unit - AOU Maggiore della Carità - Novara - Italy
(4) Neuroradiology Unit - ASST Sette Laghi - Varese - Italy
(5) Medical Physics Unit - ASST Sette Laghi - Varese - Italy

## Slide 56: Results: KAP and Effective dose (E)
Hospital 1
w/o Clarity
Hospital 1 with Clarity
Hospital 2 Biplane
84 Gycm2
158 Gycm2
26 Gycm2
21 Gycm2
111 Gycm2
233 Gycm2
9.2 mSv
11.5 mSv
3.1 mSv
1.6 mSv
10.2 mSv
19.1 mSv
Hospital 1
w/o Clarity
Hospital 1 with Clarity
Hospital 2 Biplane
Air Kerma-Area Product
Effective dose
Clarity reduces patient exposure by about 5 times, even more on fluorography.

## Slide 57: Results: Organ doses (HT)
Hospital 1 w/o Clarity
Hospital 1 with Clarity
Hospital 2 Biplane
43
525
5
42
5
194
254
177
39
40
64
888
50
221
338
259
55
24
8
9
11
150
1
11
0.8
61
74
48
10
6

## Slide 58: Inter vendor comparison of open service call length in days

## Slide 59: Early life cycle interventions
Risk assessments
Business case
Failure mode and effects
analysis ( particularly useful in new uses and projects)

## Slide 61: Failure mode and effects analysis
FMEA
RPN = Risk priority number
Assess top risks
Apply control
Eliminate top risks

## Slide 62: Documentation: write a good report

## Slide 63: Artificial intelligence

## Slide 64: AI Generated fluoro images Courtesy Laura Gambini, trinity college Dublin, Subject of workshop

## Slide 65: “Measure what is measurable, and make measurable what is not so”
― Galileo Galilei

## Slide 66: Conclusions
Fluoroscopy and angio has the highest criticality and some of the highest occupaiti0onal and patients dose with risk of tissue effects
A good patient focused quality management system aligned with health care objectives can make real patient impact
Quality control and quality assurance are part of that
Equipment testing is only one important part of quality control
It is essential that medical physics add value to results to create patient impact
It is essential that we adaptable to new environments such as AI and we continue to assess our existing and develop new techniques
Optimizations, Audits and incidents control need physics involvement to turn the quality cycle
We need more medical physicists involved in diagnostic and fluoroscopy in many countries