## Slide 1: Performance Testing (Daily, Weekly, etc.)
Chai Hong Yeong, PhD
Chair, Medical Physics World Board, IOMP
Taylor’s University, Malaysia
Chaihong.Yeong@taylors.edu.my
Joint ICTP-IAEA Training Course on
Quality Assurance, Quality Control and Optimization of Equipment and Procedures Used in Fluoroscopically-Guided Procedures

## Slide 2: Contents
What is performance test? Why?
MRT tests
CQMP tests
Some useful tips

## Slide 3: Disclaimer
I have no conflict of interest to declare in this talk.
The main references of this talk include: IAEA HH 47 report (2022), AAPM 272 report (2022), and Malaysia Ministry of Health QC guidelines.
I would like to acknowledge my colleagues at the IAEA Task Group, Kevin Wunderle, Hendrik de Vos, Francoise Malchair, Carlo Maccia, Carlos Ubeda and Virginia Tsapaki for contributing to the contents of this talk

## Slide 4: IAEA HH47 (2022)
Main References

## Slide 5: AAPM Report 272 (2022)
https://aapm.onlinelibrary.wiley.com/doi/full/10.1002/mp.15429
Main References

## Slide 6: Ministry of Health Malaysia (2022)
Radia.moh.gov.my
Main References

## Slide 7: QC for Fluoroscopy Units
  Acceptance test: new installation, prior to clinical use.
  Commissioning: ensure the system is ready for clinical use & to establish baseline. Should consider all parameters and conditions used in clinical.
  Constancy check: routine performance tests by MRTs (basic, more frequent) & CQMPs (complex, annually).
Introduction
MRT: Medical Radiation Technologist
CQMP: Clinically Qualified Medical Physicist

## Slide 8: Performance Testing
Why?
To get calibrate all the exposure parameters, to check functional performance of X-ray equipment and radiation safety around the X-ray installation.
In other words: to obtain the optimum quality diagnostic information at the lowest radiation risk to the patient.
QA checks are necessary for every X-ray equipment.
Introduction

## Slide 9: Components of a fluoroscopy unit
(Image with permission from Stephen Balter)
Introduction

## Slide 10: QC for Fluoroscopy & Angiography Units
By MRT | Frequency | By CQMP | Frequency
1. Reproducibility of the AERC | Quarterly | 1. System check | Annually
2. Image uniformity and artefacts | Monthly/Quarterly | 2. X ray beam alignment & centering | Annually
3. Monitor display check | Quarterly | 3. X ray beam collimation | Annually
 |  | 4. Tube potential accuracy | Annually
 |  | 5. mAs linearity | Annually
 |  | 6. Reproducibility of radiation output | Annually
 |  | 7. HVL measurement | Annually
 |  | 8. Accuracy of dose indices display | Annually
 |  | 9. Image Receptor ESAK rate (for II systems) | Annually
 |  | 10. Leakage radiation | Annually
 |  | 11. Scattered radiation | Annually
 |  | 12. Monitor display check | Annually
 |  | 13. Artifact evaluation | Annually
Introduction

## Slide 11: QC for Fluoroscopy & Angiography Units
By CQMP | Frequency
14. Image Quality assessment | 
14.1 Low Contrast visibility | Annually
14.2 Threshold contrast detail detectability | Annually
14.3 High contrast (spatial) resolution | Annually
15. Patient entrance surface air kerma rate | Annually
16. Maximum air kerma rate | Annually
17. Table and table pad attenuation evaluation | Annually
18. RDSR and DICOM image data verification | Annually
19. Advanced: CBCT image quality and radiation dose evaluation | Annually
Introduction
Qualitative

## Slide 12: MRT Tests
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check

## Slide 13: Some tips before we start…
MRT tests are performed more regularly (monthly or quarterly), hence the setup should be simple and straightforward, the use of comprehensive phantoms is not practical.
20 cm PMMA is commonly used to simulate body thickness.
Alternatively, 2 mm Copper plate can be used to achieve the similar beam attenuation. It can be placed on the table or on the X ray tube exit port. Make sure the copper plate covers the entire FOV.
Use service mode or manual radiographic mode whenever possible.
Make sure you have the baseline data from the commissioning for comparison purpose.
Follow the same settings & exposure parameters used during commissioning.

## Slide 14: MRT Tests
Reproducibility of the Automatic Exposure Rate Control (AERC)
MRT Tests
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check
AERC is mandatory for all fluoroscopy systems. It adjusts the exposure parameters automatically to achieve a constant predefined image quality under different X ray beam conditions
(Image with permission from Stephen Balter)

## Slide 15: AERC or ABC
Image courtesy of Colin Martin and David Sutton
MRT Tests
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check

## Slide 16: MRT Tests
Reproducibility of the Automatic Exposure Rate Control (AERC)*
MRT Tests
Also called:
  Automatic Dose Rate Control (ADRC) (in AAPM)
  Automatic Brightness Control (ABC) (in II systems)
It is important to test all the commonly used clinical imaging protocols for this test.
Frequency: Quarterly
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check
Notes: The copper filter should be at least 10 cm 3 10 cm in size and is attached to the X ray tube diaphragm.

## Slide 17: MRT Tests
Reproducibility of the Automatic Exposure Rate Control (AERC)*
MRT Tests
Set SID = 100 cm, or other clinically relevant distance tested earlier.
Place 10 cm PMMA or 1 mm copper plate on the table, ensure it covers the whole FOV. Record the exact setup (SID, table height, FOV, etc).
Set the most frequently used clinical imaging protocol (e.g. additional filtration, focal spot size, mode of operation, use of grid, etc). They should be the same settings used during baseline. Record these settings.
Make an exposure long enough for the AERC to reach a constant value (e.g. > 5 s). Record the kV, mA, filtration, pulse width, fps, and any relevant info.
Repeat exposure for each clinically relevant magnification or FOV.
Repeat the test with two other thicknesses of the attenuator (e.g. 15 cm and 20 cm PMMA, 2 mm and 3 mm Copper plate).
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check
Notes: The copper filter should be at least 10 cm 3 10 cm in size and is attached to the X ray tube diaphragm.

## Slide 18: MRT Tests
Reproducibility of the Automatic Exposure Rate Control (AERC)*
MRT Tests
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check
Copper plate can also be placed here
Notes: The copper filter should be at least 10 cm 3 10 cm in size and is attached to the X ray tube diaphragm.

## Slide 19: MRT Tests
Reproducibility of the Automatic Exposure Rate Control (AERC)
MRT Tests
Analysis:
For each setting and protocol, compare the readings with the established baseline.
Tolerances:
±5% for kV, ±20% for mA
The other parameters must be the same.
Corrective actions:
Check the geometry setup and exposure parameters. Repeat the test after resetting.
Ask CQMP to perform a more thorough investigation if the deviation persists.
If the error is reproducible, contact service engineer.
The system should not be operated without an appropriately working AERC.
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check

## Slide 20: MRT Tests
MRT Tests
2. Image Uniformity and Artefacts
Procedure:
Take an image and review it carefully on the display for uniformity and presence of any artefacts (dirt, blood, contrast agent on image receptor or other surface in the X ray beam path, table or monitor, unexpected object overlaid on anatomy).
Analysis:
No visible artefacts or grossly inhomogeneous areas should be observable.
Frequency:
Monthly / Quarterly
Example of artefact: smaller FOVs are visible on the image.
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check
(Images courtesy of Francoise Malchair)

## Slide 21: MRT Tests
MRT Tests
3. Monitor Display Check
Recommended test patterns for LED monitors:
  AAPM TG18 QC
  IEC modified TG18-OIQ
  AAPM TG270 sQC
TG-18 QC
TG-18 OIQ
TG-270 sQC
The SMPTE test pattern is grayscale insensitive, and the other features were designed to test properties of CRT displays, and are not relevant for LCD and OLED display.
Reproducibility of the AERC
Image uniformity & artefacts
Monitor display check
Notes: The SMPTE pattern (formally the SMPTE Medical Diagnostic Imaging Test Pattern) has a long history of use for film printer and display QA. It provides the user with both 0%/5% and 95%/100% contrast objects, allowing the user to qualitatively evaluate the luminance response at the top and bottom of the grayscale range, but not between. On an 8-bit system, 5% contrast corresponds to approximately 13 gray levels. The 5% contrast can always be visualized, except on systems that either deviate significantly from the DICOM GSDF, or are in environments with poor ambient light conditions. The other features on the SMPTE pattern (geometric lines, scanning artifact windows, etc.) were designed to test properties of CRT displays, and are not relevant for modern LCD and OLED display technology. As a result of the pattern’s grayscale insensitivity and CRT-specific features, this report considers the SMPTE test pattern deprecated for qualitative display evaluation in favor of either quantitative measurement or updated test patterns.

## Slide 22: CQMP Tests
Performance (Accuracy & Consistency)
Image Quality
Radiation Safety

## Slide 23: CQMP Tests
X-ray Generator related
kVp accuracy & reproducibility
mAs linearity
HVL measurement
Radiation output accuracy & reproducibility
Fluoroscopic Alarm indicator (follow local regulation)
X-ray Field related
X ray beam alignment & centering
X ray beam collimation
Radiation dose related
Image receptor ESAK rate (for II)
Typical patient ESAK rate
Patient entrance dose per frame
Maximum Air Kerma rate
PKA display accuracy
Table & pad transmission
Image Quality related
Low contrast visibility
High contrast (limiting spatial) resolution
Threshold contrast detail detectability
Artefact evaluation
Shielding related
Leakage radiation
Scattered radiation
Others
Unit assembly evaluation
Monitor display check
RDSR accuracy
CBCT
(not standardised, yet…)
CQMP Tests

## Slide 24: Before we start….
MP, you will need more comprehensive test objects and calibrated radiation dosimeters for the tests.
A multifunction detector (that can measure kV, HVL, exposure time, air kerma, air kerma rate, etc) will be helpful.
For certain tests, it would be helpful to position the C-arm at 180o (X-ray tube on top, image receptor at the bottom). However, some systems may not allow this configuration. Then, we need some creativity, e.g. 900 and get the couch to support the phantom/test objects.
If you’re using an ionization chamber, make sure that it is placed free-in-air, and at least 20 cm away from the image receptor.
If you’re using a solid-sate detector, make sure that the sensitive area facing the incident X-ray beam.
CQMP Tests

## Slide 25: Before we start….
There are typically no light beams in angiography systems. You may need to check the positioning of the test objects using fluoroscopic images (the most convenient way). But, bear in mind that you will need to be in the room when fluoroscopy is on (same role as an interventional radiologist), and PPE is important! (at least a lead apron)
Some detector readings need to be corrected for the beam energy, filtration, angulation, etc. – make sure you have the info (calibration factors) on hand. Non-pressurized ion chambers will need to be corrected for temperature and pressure too.
Some tests require high dose/dose rate to perform the test, remember to protect the image receptor! (cover with sufficient lead)
CQMP Tests

## Slide 26: Before we start….
There is no fixed sequence for the tests, however it would be practical to group certain tests together (it requires experiences). You’ll get better and quicker by practice.
There are rooms for creativity/innovation, but they must be based on fundamental radiation physics principles.
Most often, the real challenges are:
  Access to service mode/engineering mode
  To download/transfer the raw images/test results (cyber-security issue)
CQMP Tests

## Slide 27: CQMP Tests
1. Unit Assembly Evaluation
To ensure proper operation of all the mechanical and functional components of the system.
Example Check list:
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR

## Slide 28: CQMP Tests
2. X ray Beam Alignment and Centering
Objective:
To ensure that the X ray field is appropriately sized and aligned with the image receptor.
The central ray of the X ray beam should be aligned with the centre of the image receptor.
Procedures:
If possible, rotate the C arm in 1800 direction (X ray tube on top, image receptor at the bottom). Then place the alignment test tool directly on the image receptor.
If this orientation is not possible, then place the test tool on the table, with the plate as close to the image receptor as possible on top of the cylinder and level (refer figure).
If the test tool is not directly placed on the image receptor, measure the distance of the test tool from the focal spot and calculate the magnification factor (or use an object with known size, e.g. coin).
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR

## Slide 29: CQMP Tests
2. X ray Beam Alignment and Centering
Procedures:
Choose the largest possible field size and fully open the collimators.
Under fluoroscopic guidance, move the table so that the points within the cylinder are aligned (overlapping), this represents the central of the X ray beam.
Collimation test tool
Alignment test tool
Overlapped Points
(Images courtesy of Kevin Wunderle)
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR

## Slide 30: CQMP Tests
2. X ray Beam Alignment and Centering
Analysis:
Determine the distance between the central ray and the middle of the image receptor in the lateral and vertical image directions.
Tolerances:
The deviation of the X ray field and image receptor
area should not exceed 2% of the SID (eg. ±2 mm).
Frequency: Annually
Corrective actions:
Check the gantry primary and secondary angles to ensure the system is at 00 in both rotational directions. Also ensure that the procedure table and test tool are level.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR

## Slide 31: CQMP Tests
3. X ray Beam Collimation
Objective:
To verify that the X ray field is appropriately limited and does not extend beyond the image receptor assembly or the intended viewable area.
Methods:
There are multiple methods to perform this test:
If a CR system is available at the facility, use the CR plate to acquire the image.
If CR system is not available, use self-developing films, such as radiochromic film.
Or use the fluorescent screen with test object/markings. Note the staff needs to be in the room (with lead apron) performing live fluoroscopy.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR

## Slide 32: CQMP Tests
3. X ray Beam Collimation
Using (a) CR plate or (b) radiochromic film:
Rotate the C arm to 1800 (if possible) or 90o.
Place a lead plate on the image receptor, ensuring that the image receptor is fully covered by the lead in the largest FOV.
Place the CR plate or radiochromic film on the lead plate.
Use a radio-opaque ruler or other object of known size to determine the magnification factor.
Starting with the smallest FOV, irradiate the CR plate or film in acquisition mode (cine or DSA).
Repeat for each available FOV.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR

## Slide 33: CQMP Tests
3. X ray Beam Collimation
Using radiochromic film:
Lead plate placed on the image receptor
Radiochromic film
Radiochromic film with radio-opaque markings
Radiochromic film is cut and placed at the 4 quadrants of the image receptor (ensure that the film coverage exceeds the largest FOV
(Images courtesy of Francoise Malchair)
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR

## Slide 34: CQMP Tests
3. X ray Beam Collimation
FOV 23 cm
Using fluorescent screen with a test object:
(Source: IAEA HH47)
Beam collimation test tool
FOV 16 cm
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: (a) Set-up used to verify the collimation of the X ray beam using a fluorescent screen. (b) The iris collimator is visible on the edge of the image, indicating that the beam confinement is acceptable.

## Slide 35: CQMP Tests
3. X ray Beam Collimation
Analysis:
Evaluate the displayed field size for every FOV. Determine whether there is a deviation from the nominal FOV. (note: some vendors report/display the diagonal dimension of the image receptor, while others use the single side length of the square).
Tolerance:
The X ray field size should not deviate from the nominal area by more than ±2% of the SID (e.g. ±2 mm if SID is 100 cm).
Frequency: Annually
Corrective actions:
If the tolerances are exceeded, initiate engineering service.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR

## Slide 36: CQMP Tests
4. Tube Potential (kVp) Accuracy
Objective:
To verify the accuracy of the kVp selected from the operator’s console corresponding to the actual value.
Procedures:
Use service mode/manual radiographic mode to select the exposure parameters. Otherwise, use varying copper or PMMA to achieve the desired kVp and mA.
Set 4 or more clinically relevant kVp (e.g. 60, 80, 90, 110 kVp) and low mA (e.g. 2 mA).
Place the kV meter on a stable surface with its sensitive area facing the X ray tube. Ensure that the X ray beam is perpendicular and centred on the detector.
Collimate the X ray field to cover the sensitive volume of the detector. Check its position through a fluoroscopic image.
Perform exposure for every chosen kVp for over 10 s and register the results.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: (a) Set-up used to verify the collimation of the X ray beam using a fluorescent screen. (b) The iris collimator is visible on the edge of the image, indicating that the beam confinement is acceptable.

## Slide 37: CQMP Tests
4. Tube Potential (kVp) Accuracy
Analysis:
Calculate the deviation of measured kVp from the nominal ones.
Tolerance:
Maximum deviation ±5% or ±5 kVp, whichever is greater.
Frequency: Annually
Corrective actions:
If the tolerances are exceeded, first repeat the test. If the problem persists, contact service engineer.
(Images courtesy of Kevin Wunderle)
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: (a) Set-up used to verify the collimation of the X ray beam using a fluorescent screen. (b) The iris collimator is visible on the edge of the image, indicating that the beam confinement is acceptable.

## Slide 38: CQMP Tests
5. mAs Linearity
Objective:
Radiation output is defined as mGy/mAs at a fixed X ray beam spectrum (kVp, filtration) and fixed distance. This ratio should remain constant over a range of mAs.
Procedures:
Same setup as Test (4) kVp accuracy.
Place multifunction meter within the X ray field, free-in-air, if possible.
Collimate the X ray field as small as reasonably achievable, while fully cover the detector.
Set a fixed clinically used kVp (e.g. 80 kVp) and Cu filtration (if selectable).
Set the mAs on the low end of clinical use (e.g. 20 mA and 100 ms, or 2 mAs).
Make an exposure and record the mAs and air kerma (mGy).
Repeat the test for at least two additional higher mAs values used clinically.
Note: For system with dynamic filtration, it is important that the filtration remain constant throughout the test. Also, the same focal spot setting should be used for all measurements, otherwise the test may not yield appropriate results.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 39: CQMP Tests
5. mAs Linearity
(Images courtesy of Kevin Wunderle)
Analysis:
Calculate radiation output (mGy/mAs) and plot mGy/mAs versus mAs.
Tolerance:
mGy/mAs at the various mAs should all be within ±10% of each other (or coefficient of linearity ≤10%).
Frequency: Annually
Corrective actions:
First, repeat the test. If problem persists, contact service engineer.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: For mobile fluoroscopes that do not have manually adjustable mAs but use an ADRC system to determine mAs for single-shot radiographic acquisitions, the tube current/mAs linearity should be assessed as follows:
Place a minimal amount of an attenuating medium (e.g., 2–5 cm of PMMA) on the image receptor.
∙ Orient the measuring device free-in-air around the primary beam axis of the C-arm gantry.
∙ Take a radiographic exposure.
∙ Increase the PMMAthickness and repeat the previous steps, ensuring that the kV is constant. Tube voltage
is likely to remain constant for acquisitions as long as fluoroscopy is not engaged.
∙ Plot the mGy/mAs and ensure linearity to within ±10%.

## Slide 40: CQMP Tests
6. Radiation Output (kV, exposure time, air kerma) Reproducibility
Objective:
To verify that the reproducibility of radiation output are within the acceptable level.
Procedures:
Same setup as Test (4) kVp accuracy.
Select kVp from 50-120 by a step of 10, expose (fluoro then single shot) and record the results.
Calculate the reproducibility of kV, exposure time (for single shot exposure), and kerma rate or air kerma.
At 80 kVp, repeat exposure 4 times and record the results.
Analysis:
Calculate the deviation of measured parameter from the nominal ones.
Calculate coefficient of variability (COV) for the repeated measurements.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 41: CQMP Tests
6. Radiation Output Reproducibility
Analysis:
Calculate the output with the following equation:
Y = radiation output (mGy/mAs)
Mc = corrected reading of the instrument (µGy)
Q = the charge or current–time product (mAs)
dFDD = focal spot to instrument distance (m)
dref = reference distance (m)
Tolerance:
COV ≤5%
Radiation output for 80 kV, 2.5 mmAl should be in the range of 25-80 μGy/mAs. Deviations >20% from baseline as remedial level, >50% as suspension level.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 42: CQMP Tests
7. HVL Measurement
Objective:
To assess the X ray beam quality to ensure accuracy and that the X ray beam is sufficient.
Procedures:
Method A: Direct measurement
With modern multifunction meters, HVL can be obtained
simultaneously during Test (6).
Method B: If HVL cannot be measured directly
Place the radiation detector at 100 cm SDD.
Make an exposure in the radiographic mode (e.g. 70 kVp, 2 mAs).
Record the measured air kerma.
Obtain a series of dose measurement by adding thickness of Al filter until the dose reduced to 50% or less than that obtained without any added filtration.
Repeat the measurement using a higher kVp (e.g. 90 kVp, 110 kVp).
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 43: CQMP Tests
7. HVL Measurement
Method C: When kVp cannot be set manually
Use fluoroscopy mode.
Place the radiation detector within the FOV.
Place several 1 mm Al filter between the detector and the image receptor. The sheets need to be big enough to span the X ray field and enough of them achieve at least 80 kVp.
Make exposure and record the kVp and mA.
Record the dose rate for this setup (“0” mm Al).
Remove a single sheet of Al and place it in between the X-ray tube and the detector. (it can be placed directly on the surface of the X-ray tube)
Record dose rate and thickness of Al added in mm.
Repeat the steps until the dose rate is reduced to less than half of the original value. Ensure that the ABC holds a constant mA and kV.
Plot dose rate (mR/s) versus mmAl (mm) and determine the HVL from the graph.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 44: HVL Measurement
Table
Image receptor
X ray tube
Several 1 mm Al filter
Radiation detector
Holder
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests

## Slide 45: HVL Measurement
Table
Image receptor
X ray tube
Several 1 mm Al filter
Radiation detector
Holder
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests

## Slide 46: HVL Measurement
Table
Image receptor
X ray tube
Several 1 mm Al filter
Radiation detector
Holder
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests

## Slide 47: HVL Measurement
Table
Image receptor
X ray tube
Several 1 mm Al filter
Radiation detector
Holder
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests

## Slide 48: HVL Measurement
Table
Image receptor
X ray tube
Several 1 mm Al filter
Radiation detector
Holder
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests

## Slide 49: HVL Measurement
Table
Image receptor
X ray tube
Several 1 mm Al filter
Radiation detector
Holder
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests

## Slide 50: CQMP Tests
7. HVL Measurement
50% transmission
HVL
Example of analysis results
Tolerance: ±10% compared to baseline
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 51: CQMP Tests
8. Accuracy of Dose (PKA) Display
Objective:
To verify the accuracy and reproducibility of dose indices display.
Procedures:
Place the radiation detector on the table or on the image receptor.
X ray field-size must be determined in the plane of air kerma measurement. This can be accomplished using radiochromic film or radio-opaque rulers to measure the field size.
If an ionization chamber is used, backscatter must be considered and both detector and film must be positioned at least 20 cm from the table on polystyrene blocks.
Kerma (or kerma rate and time) is recorded for fluoroscopy and acquisition modes.
kVp should be in the range of 90-100 kVp, repeat exposure 4 times and record the results.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 52: CQMP Tests
8. Accuracy of Dose (PKA) Display
Analysis:
If using radiochromic film, determine the field size using a ruler (or appropriate calibrated computer measuring tool).
Calculate the measured PKA by multiplying the measured air kerma and X ray beam area.
Calculate the deviation between the measured and displayed values. (Make sure that all measurements are corrected (for inverse square distance and backscatter factor) and converted to the same units used on the display)
Tolerance:
It is expected that the PKA meter’s uncertainty will not be higher than ±25%. If the system has an indicator instead of an individual PKA meter, then its accuracy should be better than ±25%
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 53: CQMP Tests
8. Accuracy of Dose (PKA) Display
Corrective actions:
If the PKA meter’s uncertainty is higher than ±25%, a correction factor should be introduced to apply to clinical radiation dose indices.
If the deviation is outside the tolerances or is inconsistent, then a new calibration may be necessary.
If the performance of other elements of the X ray unit is adequate, the unit may be used until the PKA meter is repaired.
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 54: CQMP Tests
9. Image Receptor Entrance Surface Air Kerma Rate (for II only)
Objective: To verify that the air kerma rates at the entrance of the image receptor are within the acceptable limits.
Procedures:
Place 20 cm PMMA on the table (or 2 mm Copper on the X ray tube).
Use maximum SID, or SID that is commonly used in clinical setting.
Note: If an anti-scatter grid is used clinically, make sure it is in place during the test.
Open the collimators to the size of the image receptor.
Place the radiation detector as close as possible to the image receptor, or tape it on the image receptor. Record the SDD.
Apply fluoroscopy using a clinical imaging protocol and record the air kerma rate.
Repeat the fluoroscopy for all available field sizes, dose rate settings and different imaging protocols used under normal clinical conditions.
Correct all measurements for distance using inverse square law to calculate the dose rate at the image receptor (if dosimeter is not attached to the image receptor).
CQMP Tests
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 55: CQMP Tests
9. Image Receptor Entrance Surface Air Kerma Rate (for II only)
Analysis:
Apply the appropriate correction factors to calculate the entrance surface air kerma rate at the image receptor surface (e.g. attenuating layers, anti-scatter grid, table, AERC sensor, etc.). The transmission factors are usually found in the user’s manual.
Tolerance:
The typical image receptor entrance surface air kerma rate in fluoroscopy mode is between 0.2 and 1 µGy/s for an applied tube voltage of 70 - 80 kVp, after grid attenuation.
The deviation from the established baseline values should not be more than ±25% (remedial level), and ±50% (suspension level).
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 56: CQMP Tests
10. Patient Entrance Surface Air Kerma (ESAK) Rate
Objective: To verify that the ESAK rate is under the limits for the different imaging protocols, FOV & operation modes used under normal clinical conditions. The measured results are also indicative for the evaluation of expectable peak skin dose levels.
Procedures:
Expose the PMMA phantom using a clinical imaging setting, record the dosimeter reading, kV, mA and the image receptor settings - repeat two more times.
Repeat the measurement for a range of clinically used FOVs, dose rate settings and differing imaging protocols used under normal clinical conditions.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 57: CQMP Tests
10. Patient Entrance Surface Air Kerma (ESAK) Rate
Analysis:
Correct the readings with appropriate correction factors (e.g. pressure, temperature, calibration coefficient, etc.) for the detector used (if applicable).
Correct the measured ESAK rates to the reference point using the inverse square law. When performing the evaluation, keep in mind that ESAK includes backscatter, while incident air kerma does not.
For an angiography system, the patient entrance reference point is set by the manufacturer and it mainly depends on the type of the system according to the IEC standard.
Note: Instead of measuring the values, the displayed air kerma rates could be taken, if appropriate corrections are applied.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 58: CQMP Tests
10. Patient Entrance Surface Air Kerma (ESAK) Rate
Tolerance:
Follow national regulations. In case there are no national regulations, following values can be applied (taken from various international standards):
For fluoroscopy modes, with 20 cm PMMA:
Normal mode ≤ 50 mGy/min
High dose rate ≤ 100 mGy/min
For any thickness of patient, with PMMA:
Normal mode ≤ 100 mGy/min
High dose rate ≤ 200 mGy/min
For Cine or acquisition modes, with 20 cm PMMA:
Variation ≤ 20% with respect to the reference values.
0.03 - 0.12 mGy/image for interventional cardiology.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 59: CQMP Tests
11. Maximum Air Kerma Rate
Objective:
To evaluate the maximum air kerma rates of a fluoroscopy equipment. The radiation output levels are typically limited by local or national regulations. This test may be modified to comply with those regulations.
Procedures:
Center the radiation detector within the X ray field, free-in-air.
If possible, the detector should be placed in the plane that the air kerma rate is to be determined; otherwise, the measurement must be corrected to that plane.
For C arm, the maximum air kerma rate is determined at 30 cm from the face of the image receptor, at any SID.
If the system has variable SID, this test should be performed at the min and max SID to ensure that the output properly tracks with the changing SID.
Collimate the X ray field as small as reasonably achievable, but cover the entire detector.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 60: CQMP Tests
11. Maximum Air Kerma Rate
Procedures:
Place sufficient attenuator in the X ray beam to drive the output to the maximum rate. The attenuators should be placed at least 10 cm from the detector toward the image receptor.
Record the technique (kVp and mA) ensuring that they are the maximum for the system, and record the measured air kerma rates.
Maximum air kerma rate measurements should be made for different fluoroscopic system settings:
  All fluoroscopic dose curves (there are typically 2-3).
  Largest and smallest FOV.
  Highest pulse rate (or continuous if system has this option) and at least 1 other clinically used pulse rate.
Repeat the test for a commonly used clinical FOV and frame rate for at least one clinically used acquisition.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 61: CQMP Tests
11. Maximum Air Kerma Rate
Tolerance:
Follow local regulations.
If local regulations not available:
  Fluoroscopy Normal Dose Rate Limit = 88 mGy/min
  Fluoroscopy High Dose Rate Limit = 176 mGy/min
  Acquisition mode = currently has no applicable dose rate limit (NOTE: air kerma rates measured under this geometry may exceed 2-3 Gy/min. Caution should be used when allowing dose rates to reach these levels and consider lowering the maximum achievable rates).
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 62: CQMP Tests
12. Table and Table Pad Attenuation
Objective: To evaluate the attenuation properties of the procedure table and patient mattress. The manufacturer must specify the maximum value of the attenuation equivalence, according to the IEC 60601-2-42 standards.
Procedures:
Using a PA projection, set the table height to approximately isocentre.
Move away the table and pad.
Set the X ray field size to 10 cm x 10 cm.
Centre the radiation detector within the X ray field, free-in-air (i.e. no table and pad).
Protect the image receptor with a lead plate.
Make a radiographic exposure at ~100 kVp and ~ 1 mAs, record the air kerma.
Move the table and mattress back to the field.
Make another radiographic exposure with the exact same settings, record the incident air kerma.
Note: If radiographic exposure cannot be done, use fluoroscopic mode and record the air kerma rate.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 63: CQMP Tests
12. Table and Table Pad Attenuation
Analysis:
Tolerance:
In general, the transmission should be 0.6 - 0.8 at 100 kVp.
If the transmission factor < 0.6 at 100 kVp, the attenuation properties of the table and mattress should be investigated.
Only tables and mattresses intended for use in X ray procedures should be used, others may overly attenuate the X ray beam.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 64: CQMP Tests
13. Artefact Evaluation
Same as the MRT test
A grid patterns may be used to evaluate pincushion and S distortion on II systems
S distortion from II
Blood infiltrating collimator assembly
Grid artefact
Gassing artefact from an II
(Images courtesy of Hendrick de Vos)
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 65: CQMP Tests
14. Monitor Display Check
Same as the MRT test
TG-18 QC
TG-18 OIQ
TG-270 sQC
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 66: CQMP Tests
15. Image Quality Assessment
15.1 Low Contrast Visibility
Tolerance:
According to the test object used, the tolerance from the baseline should be decided and proposed by the CQMP in cooperation with the FGIP physician & MRT.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 67: CQMP Tests
15. Image Quality Assessment
15.2 Threshold Contrast Detail Detectability
Tolerance: HT or IQF should not exceed 10% difference from the baseline.
n = number of detail diameter groups,
HT(A) = threshold detection index calculated from the image.
HTref(A) = threshold detection index calculated from a reference image of a system known to be in good adjustment.
Kref = Input air kerma rate for the reference iagme (in nGy/s).
Km = Measured input air kerma rate of the system (in nGy/s).
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 68: CQMP Tests
15. Image Quality Assessment
15.3 High Contrast (Spatial) Resolution
Tolerance:
There should not be any change in limiting spatial resolution from the baseline.
Examples of TOR 18FG and Huttner resolution test pattern.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 69: CQMP Tests
16. Leakage Radiation
Objective:
To evaluate the amount of radiation leaking from the X ray tube or collimator assembly.
Procedures:
Close the collimator blades completely or to the smallest size available.
Place lead sheet over the X ray tube exit to protect the image receptor.
Select a voltage about 10 kVp lower than the maximum kVp (e.g. 110 kV) and a suitable mA and exposure time (e.g. 200 mA and 0.1 second) and make an exposure..
Position the survey meter at 1 m from the X ray tube focal spot.
Perform measurements at multiple positions and record the maximum air kerma rate measured for each position. The amount of measurement positions can be determined considering the directionality of the survey meter used but should at least include the anode and cathode sides of the X ray tube and consider likely positions where personnel may stand in relation to the tube assembly.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 70: CQMP Tests
16. Leakage Radiation
Procedures:
Note:
If the system can only use fluoroscopy mode, use the ABC mode for 5 s at the highest kV.
Leakage measurements should always be taken at the maximum practical kVp, however, caution should be exercised particularly where tube voltage errors may be apparent. If there is uncertainty, use a tube voltage not more than 10 kVp less than the maximum.
Tolerance:
The maximum air kerma rate should be <1 mGy/h at 1 m from the focal spot.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 71: CQMP Tests
17. Scattered Radiation
Objective: To verify that the amount of scattered radiation from the fluoroscopy system at every adjacent occupational area are within acceptable level.
Procedures:
Place 20 cm PMMA on the table and position it at the centre of the X-ray field.
Open collimator fully and ensure that it cover the entire size of the phantom.
Cover the image receptor with lead.
Position the survey meter at the area of interest and make an exposure with the highest exposure factor clinically used.
Tolerance:
±50% from the baseline value.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 72: CQMP Tests
18. RDSR and DICOM image data verification
Objective:
To verify the accuracy of radiation dose structured report (RDSR) and DICOM image data.
Procedures:
Expose any phantom using the most used fluoroscopy mode and cine mode.
Make several exposure while noting the technical and geometric parameters.
Send the images and RDSR to the PACS and/or DACS, if they are not automatically configured.
Analysis:
Check the image DICOM header and RDSR for consistency and accuracy. Compare technical and geometric information from that recorded earlier. Check the RDSR in DICOM format for completeness.
Tolerances:
DICOM headers and values in RDSR must be identical. All mandatory fields in RDSR must be complete and accurate, table position is sometimes incorrect.
Unit Assembly
Beam alignment & centering
Beam collimation
kVp accuracy
mAs linearity
Radiation output
HVL
PKA accuracy
Image Receptor ESAK rate (II only)
Patient ESAK rate
Max Air Kerma rate
Table and pad
Artefact
Monitor Display
IQ (qualitative)
Leakage
Scattered
RDSR
CQMP Tests
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 73: CQMP Tests
19. CBCT Image Quality and Dose Evaluation
Advanced
EFOMP-IAEA-ESTRO joint working group and AAPM have recently released the QC protocols for CBCT.
The EFOMP-IAEA-ESTRO protocol focuses on image quality and radiation output, where radiation dose is described by the EURADOS WG12 in Annex 6 as per a review of the literature.
The AAPM report reviews the image acquisition and reconstruction characteristics of C arm CBCT and provides guidance on QC of the systems.
Dose measurements in CBCT are challenging and may be associated with many limitations, like the size and positioning of the phantom, the non-uniform distribution of the beam, etc. Measure of radiation output reproducibility is suggested using one of the many available methods (e.g. CTDI, PKA, image receptor dose, Exposure Index, etc.).
Notes: This test is a surrogate for evaluating X ray tube current accuracy. X ray tube current can only be directly measured using invasive techniques requiring access to the high voltage generator power supply, which is quite dangerous and should not be attempted by anyone other than qualified service engineers. However, the mAs linearity test provides a means to demonstrate that changes in the set mAs result in a relatively constant mGy/mAs. This allows one to conclude that the tube current is responding to change as expected.
For systems where independent control of the technical parameters is possible:
  Place lead or other highly attenuating material within the X ray beam, sufficient to fully protect the image receptor when making exposures. Take care not to place this material too close to the measurement device, a minimum distance of 10 cm should separate the attenuation material and the measurement device. This is especially important if an ionization chamber is used to make these measurements.

## Slide 74: Summary
Test | Acceptance level | Remedial level | Suspension Level
1. Unit Assembly | All mechanical and safety parts function properly |  | 
2. X ray beam alignment & centering | ≤ 2% of SID | > 2% of SID | > 3% of SID
3. X ray beam collimation | Total deviation <4% of SID | ≥ 4% of SID Radiation area >1.1x image receptor | > 5% of SID Radiation area >1.25x image receptor
4. kVp accuracy | Max. deviation ≤ ±5 kV or ±5%, whichever greater | Max. deviation > ±5 kV or ±5%, whichever greater | Max. deviation > ±10 kV or ±10%, whichever greater
5. mAs linearity | ≤ 10% | > 10% | > 20%
6. Reproducibility of radiation output | CoV ≤ 5% | CoV > 5% | CoV > 10%
7. HVL | Depending on kVp and mmAl, refer the chart |  | 
8. Accuracy of PKA | ≤±25% |  | 
Example. Follow local/national regulation (if N/A, then refer international guidelines)

## Slide 75: Summary
Test | Acceptance level | Remedial level | Suspension Level
9. Image Receptor ESAK rate (II) |  |  | 
i. 11 cm to < 14 cm | ≤ 120 μGy/min | > 120 μGy/min | > 180 μGy/min
ii. 14 cm to < 23 cm | ≤ 90 μGy/min | > 90 μGy/min | > 135 μGy/min
iii. ≥ 23 cm | ≤ 60 μGy/min | > 60 μGy/min | > 90 μGy/min
10. Typical Patient ESAK Rate |  |  | 
With AERC | ≤ 25 mGy/min (2.5 R/min) | NA | ≥ 25 mGy/min (2.5 R/min)
Without AERC | ≤ 12 mGy/min (1.2 R/min) | NA | ≥ 12 mGy/min (1.2 R/min)
10.2 Patient Entrance Dose Per Frame |  |  | 
(a) Normal Digital Fluoro mode | ≤ 2 mGy/frame | NA | > 2 mGy/frame
(b) Cardiac Mode | ≤ 0.2 mGy/frame | NA | > 0.2 mGy/frame

## Slide 76: Summary
Test | Acceptance level | Remedial level | Suspension Level
11. Max Air Kerma Rate | Normal Dose Rate = 88 mGy/min High Dose Rate = 176 mGy/min |  | 
12. Table/Pad transmission | 0.6 - 0.8 at 100 kVp | <0.6 | 
13. Artefact evaluation | No artifact |  | 
14. Monitor display check | Pass all the visual check |  | 
15. Image Quality assessment | Refer test objects manual |  | 
15.1 Low Contrast visibility |  |  | 
15.2 TCDD |  |  | 
15.3 High contrast (spatial) resolution |  |  | 
16. Leakage radiation | ≤ 0.1 mGy or ≤ 10 mR | ≤ 1 mGy (≤ 100 mR) | > 1 mGy or > 100 mR
17. Scattered radiation | ≤ 0.1 mGy or ≤10 mR | NA | > 0.1 mGy or > 10 mR
18. RDSR and DICOM data verification | Must be accurate |  | 

## Slide 77: Summary
Regulatory Guidelines
QC Report (example only)
(Courtesy of Medical Physics Services Sdn Bhd)
3.5 Repeatability of Radiation Output (Radiographic Mode)
3.6 Coefficient of Linearity

## Slide 78: Where to Get More Information
2022

## Slide 79
IAEA Task Group (2022-202X)
“Guidance related to Quality Assurance and optimization in fluoroscopically-guided interventional radiology”

## Slide 80: Thank You!
Chai Hong Yeong, PhD
Chair, Medical Physics World Board, IOMP
Taylor’s University, Malaysia
Chaihong.Yeong@taylors.edu.my
Joint ICTP-IAEA Training Course on
Quality Assurance, Quality Control and Optimization of Equipment and Procedures Used in Fluoroscopically-Guided Procedures