## Sheet1
TASKS FOR USING QC DATA AND CALCULATING QC PARAMETERS [ FOR RADIOGRAPHY ]
| DATA: |  |  |  | X-RAY GENERATOR AND TUBE MEASUREMENTS
These measurements were made with a keithley kVp divider s/n xxxxx, digital storage oscilloscope
Gould 450 s/n zzzzz, and an MDH electrometer s/n yyyy with 6 cc chamber .
[8 empty] | FDD(cm)= | 100
| Focus | Set kV | Set mA | Set Time | Set mAs | Meas kV | Meas T | Meas exp | Air kerma
|  | (kV) | (mA) | (ms) | (mAs) | (kV) | (ms) | (mR) | (mGy)
| B | 60 | 200 | 100 | 20 | 62 | 101 | 59 | 0.51271
| B | 80 | 200 | 100 | 20 | 83 | 105 | 98.5 | 0.85596
| B | 100 | 200 | 100 | 20 | 104 | 103 | 151 | 1.3122
| B | 120 | 200 | 100 | 20 | 123 | 103 | 224 | 1.9466
| F | 50 | 100 | 200 | 20 | 50 | 202 | 46 | 0.39974
| F | 70 | 100 | 200 | 20 | 73 | 205 | 101 | 0.87769
| F | 90 | 100 | 200 | 20 | 88 | 202 | 167 | 1.4512
| F | 110 | 100 | 200 | 20 | 111 | 204 | 221 | 1.9205
| B | 80 | 200 | 100 | 20 | 82 | 105 | 104 | 0.90376
| B | 80 | 200 | 100 | 20 | 83 | 105 | 106 | 0.92114
| B | 80 | 200 | 100 | 20 | 83 | 105 | 105 | 0.91245
| B | 80 | 25 | 100 | 2.5 | 83 | 107 | 18.4 | 0.1599
| B | 80 | 300 | 100 | 30 | 84.7 | 105 | 198 | 1.7206
| B | 80 | 500 | 100 | 50 | 90 | 103 | 305 | 2.6505
| B | 80 | 200 | 20 | 4 | 83 | 19 | 20.5 | 0.17814
| B | 80 | 200 | 400 | 80 | 83 | 411 | 440 | 3.8236
| B | 80 | 200 | 800 | 160 | 83 | 780 | 814 | 7.0737
Half Value Layer Measurements are shown below
| B+0mm Al | 80 | 200 | 100 | 20 |  |  | 114 | 0.99066
| B+0mm Al | 80 | 200 | 100 | 20 |  |  | 110 | 0.9559
| B+1mm Al | 80 | 200 | 100 | 20 |  |  | 83.5 | 0.72562
| B+2mm Al | 80 | 200 | 100 | 20 |  |  | 67.1 | 0.5831
| B+3mm Al | 80 | 200 | 100 | 20 |  |  | 50.9 | 0.44232
| B+4mm Al | 80 | 200 | 100 | 20 |  |  |  | #VALUE!
CALCULATE THE FOLLOWING PARAMETRES USING THE DATA FROM THE TABLE
1 | X-ray beam filtration (HVL) - plot the graph
2 | Dose Output Consistency % [ 100*(st.dev)/(average) ] for BF
3 | kVp Consistency % [ 100*(st.dev)/(average) ] for BF
4 | Timer Consistency % [ 100*(st.dev)/(average) ] for BF
5 | kVp accuracy [ 100*(mean error)/(real value) ] for BF and FF
6 | kVp Linearity [ plot the graph mGy/mAs against kVp2 ] for BF and FF
7 | Timer accuracy [ 100*(mean error)/(real value) ] for BF
8 | Dose Output variation with mA [ 100*st.dev/average ] for BF
9 | kVp variation with mA [ 100*st.dev/average ] for BF
10 | TASK:
Calculate with how many kV will drop the HV, if the Exposure is 100 kV and 100 mA
USING PARAMETERS:
X-ray generator high voltage (HV) Transformer ratio ~ 1:500
Hospital voltage input from main - 200 V
Hospital main cables resistance - 0.2 ohms