Laboratory 1

Assessment Spatial Resolution and Contrast Resolution from Images taken from Digital Fluoroscopic equipment  and comparison with an old Analogue Fluoroscopic Equipment

The material includes the printout of the EMERALD Task 8.1 (Fluoroscopic Quality Control) and additional pages with Test Object passport data and real images from Digital fluoroscopic equipment (Image Intensifier + Digitizer) plus an Analogue equipment image (final page). It is advisable the students to work on groups of 2-3 people, and agree on results in their groups.

A. Assessment of Limiting Spatial Resolution (Lp/mm at 10% contrast)

- Using the Huttner test object data from the 5th page, identify the spatial resolution (Line pairs per mm) of each group of objects.

- Assess visually the following 4 pages (6th, 7th, 8th, 9th) and write down their limiting spatial resolutions.

- Compare the results from these 4 pages and  comment on these (how the images have been obtained; are these suitable for assessment; what is the difference between them).

B. Assessment of Contrast Resolution and Contrast Detail diagram

- Using the Leeds TO10 test object data from the 11th page, identify the contrasts and sizes of each object.

- Assess visually the image on page 12 and write down your results for each row (from A to M) – you may use the first two empty columns from the upper table at page 10.

- Using Excel software draw the Contrast – Detail curve of this image (the size of each row on X  against its limiting contrast on Y ).

- Assess visually the image on page 13 and write down your results for each row (from A to M) – you may use the second two empty columns from the upper table at page 10.

- Using Excel software draw the Contrast – Detail curve of this image (the size of each row on X  against its limiting contrast on Y ).

- Plot both diagrams in one graph, compare the results from these two Images and  comment on these (which equipment produces better C/D, why we have image deterioration, etc.).

- Knowing that the image on page 12 is from a new Digital Fluoroscopic Equipment, and the image on page 13 is from an old Analogue Fluoroscopic Equipment comment on the sources of noise in both images.

S. Tabakov