## Sheet1
Radiological Protection and Quality Assurance Report
[5 empty] | Unit type : |  | Unercouch + II
Hospital/Clinic: |  |  |  |  | Generator/Manuf.: |  | Philips
Type:
Department : |  |  |  |  | X-Tube / Type:
Ser.No.
Room/Unit: |  |  |  |  | Image Intens/Type:
Ser.No.
Date of QA survey: |  |  |  |  | Reference no.:
SUMMARY AND RECOMMENDATIONS
Report compiled by:
Dr. S.D.Tabakov
AUTOMATIC DOSE MEASUREMENT
These measurements were made with an acrylic phantom and MDH electometer with 60cc ion.chamber.
Set up distances |  |  | cm
Source to image intensifier |  |  | 100
Source to phantom |  |  | 80
Chamber to chamber |  |  | 35
Medium dose setting
Field size | Read kV | Read mA | Phantom | I.I. entrance dose |  | Phantom surf' dose |  | Grid in
cm |  |  | thick' mm | (mR/min) | (mGy/s) | (mR/min) | (mGy/min)
30 | - | - | 0 | - | - | - | -
| 75 | 0.3 | 55 | 29.9 | 0.0043305 | 264 | 2.2704
| 75 | 1.4 | 110 | 43.3 | 0.0062713 | 1010 | 8.686
| 75 | 5.6 | 165 | 68.5 | 0.0099211 | 3880 | 33.368
| - |  | 165 +0 .5Pb | - | - | - | -
23 | - | - | 0 | - | - | - | -
| 75 | 0.8 | 55 | 59.5 | 0.0086176 | 579 | 4.9794
| 75 | 3.3 | 110 | 79.9 | 0.011572 | 2270 | 19.522
| 75 | 13.8 | 165 | 124 | 0.017959 | 8560 | 73.616
| - | - | 165 +0 .5Pb | - | - | - | -
17 | - |  | 0 | - | - | - | -
| 75 | 1.4 | 55 | 90.5 | 0.013107 | 928 | 7.9808
| 75 | 5.9 | 110 | 112 | 0.016221 | 3710 | 31.906
| 90 | 10 | 165 | 138 | 0.019987 | 8370 | 71.982
| - | - | 165 +0 .5Pb | - | - | - | -
Medium dose setting
Image Intensifier entrance dose rates with 1mm of Cu in the beam.
II field | Read kV | Read mA | Dose rate
size cm |  |  | (mR/min) | (mGy/s)
30A | 75 | 0.9 | 10.9 | 0.0015787 | A=Automatic
23A | 75 | 2.4 | 24.8 | 0.0035919 | M=Manual
17A | 75 | 4.2 | 40.2 | 0.0058223
17M | 80 | 1 | 16.5 | 0.0023897
ANALOGUE IMAGE QUALITY
These measurements were performed with Leed's test objects
All measurements made in medium dose setting
GS2 with 1mm copper
Field size | Read values |  | No. of | Lead | Comment
cm | kV | mA | steps | periphery
- | - | - | 10 | Circular | Accept
M1
Field size | Read values |  | Horizontal | Vertical | Distortion
cm | kV | mA | squares | squares
- | - | - | - | - | -
- | - | - | - | - | -
- | - | - | - | - | -
Huttner
Field size | Read values |  | Group no. | Spatial
cm | kV | mA | visible | freq lp/mm
30 | 80 | 0.6 | 3 | 0.63
23 | 80 | 0.6 | 6 | 0.9
17 | 80 | 0.6 | 8 | 1.12
N3 with 1mm copper
Field size | Read values |  | Detail no. | Contrast
cm | kV | mA
- | - | - | - | -
- | - | - | - | -
- | - | - | - | -
TO10 with 1mm copper
Row | Detail size |  | Detail number and contrast with field size (using ABS)
| mm | 30cm |  | 23cm |  | 17cm
|  | 80kV | 0.2mA | 80kV | 0.6mA | 80kV | 1.2mA
A | 11.1 | 6 | 0.032 | 6 | 0.032 | 6 | 0.032
B | 7.9 | 6 | 0.032 | 6 | 0.032 | 6 | 0.032
C | 5.6 | 6 | 0.032 | 6 | 0.032 | 6 | 0.032
D | 4 | 6 | 0.045 | 6 | 0.045 | 6 | 0.045
E | 2.8 | 5 | 0.066 | 6 | 0.045 | 6 | 0.045
F | 2 | 5 | 0.066 | 5 | 0.066 | 5 | 0.066
G | 1.4 | 4 | 0.16 | 6 | 0.086 | 6 | 0.086
H | 1 | 3 | 0.23 | 4 | 0.16 | 5 | 0.123
J | 0.7 | 2 | 0.35 | 3 | 0.23 | 4 | 0.16
K | 0.5 | 1 | 0.93 | 3 | 0.5 | 4 | 0.35
L | 0.35 | 0 | 0.99 | 1 | 0.93 | 3 | 0.5
M | 0.25 | 0 | 0.99 | 0 | 0.99 | 1 | 0.93
ANALOGUE IMAGE QUALITY (cont.)
Video voltage
(TV monitor input taken with T-connector )
Test object E1(Lead edge) with 1mm Cu
| Video signals (mV) |  | Field size | II Field size
Sync. pulse/ blanking
Blanking/black level
Loss of contrast (black)
Camera noise (black) P-Pmax
Blank/white ampl.
Camera & quantum noise (white) P-P
Vignetting slope
Dose rate mR/min
kVp/mA